## Slide 1
Atomic Force Microscopy
Alexandre Dazzi, Institut de Chimie Physique
alexandre.dazzi@universite-paris-saclay.fr

## Slide 2
I.Introduction
80’s, strong willing to manipulate and observe nanometric objects
Optical Microscopy
Electron Microscopy
res ≈ 300 nm
res ≈ 0.1 nm

## Slide 3
« Feeling is better than seeing »
or the born of nanotechnology
Gerd Binnig
(Physics Nobel Prize 1986)
1985 Atomic Force Microscope
Silicon tip to palpate the surface.
Cantilever to move the tip.
No constraint on the surface nature
I. Introduction

## Slide 4
II. Tip-Surface interaction
How to describe the tip-surface interaction?
cantilever
tip
sample
Fint
Fint
sample
Rtip
II. Tip-surface interaction

## Slide 5
II. Tip-Surface interaction
A) Atom-atom interaction
Lennard-Jones potential describes the interaction between atoms or small molecules.
A,B constant related to nature of atoms and r distance between atoms
M1
M2
r
Repulsive interactions
Attractive interactions
(Van der Waals)

## Slide 6
II. Tip-Surface interaction
Lennard-Jones Potential for Si-Si
Attractive region
repulsive region

## Slide 7
II. Tip-Surface interaction
B) Tip-surface interaction
Fint
sample
Rtip
Fint
sample
Rtip
B.1) Attractives forces
Fatt
(Van der Waals)

## Slide 8
II. Tip-Surface interaction
B.1) Attractives forces for a sphere-plane interaction
R
D
H Hamaker constant, R radius of the tip,
D distance bertween tip and sample surface
Fatt

## Slide 9
II. Tip-Surface interaction
B.2) Repulsive forces for a sphere-plane interaction
B.2.1) Hertz contact
h indentation
a radius of contact area
R radius of the sphere
E Young modulus
n Poisson coefficient
R1
R2
h
Frep
Surface of contact
a

## Slide 10
II. Tip-Surface interaction
B.2) Repulsive forces for a sphere-plane interaction
B.2.2) Sphere-plane
Rtip
h
D
Etip
Esample

## Slide 11
II. Tip-Surface interaction
B.2) Repulsive forces for a sphere-plane interaction
B.2.2) Sphere-plane
How to define the distance of contact ?
D
0
dc
0
dc=0.165 nm

## Slide 12
II. Tip-Surface interaction
B.2) Repulsive forces for a sphere-plane interaction
B.2.2) Sphere-plane
Rtip
h
D
Etip
Esample
0
dc
Where D<dc
(only true in contact)
In D=dc,
but

## Slide 13
II. Tip-Surface interaction
B.3) Sphere-plane interaction force
Rtip
D
0
dc
Etip
Esample
If D>dc
If D≤dc
DMT model (Derjagin-Muller-Toporov)
Fs-p
z

## Slide 14
II. Tip-surface interaction
Si-Si

## Slide 15
III. Interaction force detection
III. Interaction force detection
A) Use of a cantilever to feel the force
F
z
x
Bending moment :
L
Euler-Bernouilli equation :
E, Young’s modulus of the cantilever
I, area moment of inertia of the cantilever cross section
w
t

## Slide 16
III. Interaction force detection
F
z
x
L
Euler-Bernouilli equation :
Boundaries conditions :

## Slide 17
III. Interaction force detection
F
z
x
L
B) Deflection measurement
Z
Deflection Z, is the flexure at the end of the cantilever.
with kc the stiffness of the cantilever
(Hooke’s law)

## Slide 18
III. Interaction force detection
Si Cantilever
10µm
Apex curvature radius 10 nm
400-500 µm
Pour le mode contact
Tapping mode
High resolution
contact | L=450 µm | k=0.1 N/m | f = 6 kHz
tapping | L=225 µm | k=40 N/m | f= 150 kHz

## Slide 19
Optical detection of the force
III. Interaction force detection
photodiode
Laser diode
cantilever

## Slide 20
III. Interaction force detection
B) Deflection measurement
HG
HD
BG
HD
+
+
-
Itop
Ibottom
Imes=Itop-Ibottom
F
photodiode
F

## Slide 21
IV. Static mode
x
PiezoZ
Dz
Fint
Z
D
A) Force curve
Fr
Equilibrium of forces :
Line with kc slope
DMT curve

## Slide 22
IV. Static mode
A
A
DzA
B
B’
DzB
C
C’
DzC
D’
D
DzD
E
DzE
D
Dz
B
B’
C
C’
D
D’
E
Force curve F(Dz)
Fs-p
piezoZ
approach
retract

## Slide 23
IV. Static mode
Force curveF(Dz)
On hard surface
Force curve F(Dz)
On soft surface

## Slide 24
IV. Static mode

## Slide 25
IV. Static mode
Application : Spectroscopy of single protein
Ubiquitin

## Slide 26
IV. Static mode
B) Imaging mode
x
PiezoZ
Dz
PiezoXY
PiezoZ actuator : surface approach
PiezoXY actuators : scan of the surface in XY plane
F
Surface scanning:
Approach
Scanning X and Y
Top view
Y
X

## Slide 27
Constant force Imaging
IV. Static mode
4 quadrants
Laser diode
piezo
sample
PI controller
setpoint
(F)
(F0)

## Slide 28
IV. Static force
A
B
F0
B
A
h
h
Dz
F
DA
DA
F0
F1
F1
DB
t
DB
F2
F2
t
t
t
F
Dz
h’=DA-DB
Modelling force constant imaging (F0)

## Slide 29
IV. Static force
Contact mode images

## Slide 30
V. Dynamic mode (tapping)
PiezoZ actuator : surface approach
PiezoXY actuators : scan of the surface in XY plane
Piezo drive : select the cantilever frequency of oscillation
x
PiezoZ
Dz
PiezoXY
Piezo drive
w
Amplitude
Signal generator
Advantage : soft and fragile sample imaging
(no more shear effect during the scan)

## Slide 31
CONTACT MODE
TAPPING MODE
AFM image
AFM image

## Slide 32
V. Dynamic mode
Euler-Bernouilli equation :
Solution :
with
and
A) Cantilever modes
frequency
damping

## Slide 33
V. Dynamic mode
Euler-Bernouilli equation :
z
x
L
Boundaries conditions :
Eigenvalues (modes) equation of the cantilever :
Mode n | bnL
0 | 1.8750
1 | 4.6941
2 | 7.8547

## Slide 34
V. Dynamic mode
Dispersion equations
z
x
L
cantilever
E, Young’s modulus of the cantilever
I, area moment of inertia of the cantilever cross section
S, cross section : S=wt
r, cantilever density
mass-spring
kc, spring constant
m, mass
Z
Z

## Slide 35
V. Dynamic mode
Each mode can be described by the mass-spring model.
Cantilever parameters :
Where mn* is the mass of the mode
Mode n | mn*
0 | 0.24267 m
1 | 0.00618 m
2 | 0.00079 m

## Slide 36
V. Dynamic mode
B) Motion equation of cantilever modes.
where
, G damping, and Zn the deflection of teh nth modes
x
Z

## Slide 37
V. Dynamic mode
C) Motion equation of cantilever driven by a piezo.
where
, G damping, and Zn the deflection of teh nth modes, Fpz force applied by the piezo
x
Z
Piezo drive
w
Signal generator

## Slide 38
V. Dynamic mode
Solution :
Solution :

## Slide 39
Exemple of 1rst mode of a cantilever
V. Dynamic mode
Amplitude(r)
Phase(j)

## Slide 40
V. Dynamic mode
D) Motion equation of cantilever driven by a piezo and in interaction with the surface.
where
, G damping, and Zn the deflection of teh nth modes, Fpz force applied by the piezo
x
Z
Piezo drive
w
Signal generator
piezoZ
Fint
Dz
non linear term

## Slide 41
V. Dynamic mode
Principle of least action (or principle of stationary action)
L is the Lagrangian of the mechanical system and defined by :
where T is the kinetic energy and U potential energy
S is the action defined as :
where t1 and t2 are 2 different times corresponding at 2 different positions of z.

## Slide 42
V. Dynamic mode
The principle of least action is :
where Xi is the variables of the system
For a conservative system the equation of motion is found by :

## Slide 43
V. Dynamic mode
Principle of least action in our case can be expressed as :
Stationary solution can be written like :
Kinetic energy
potential energy
of the spring
potential energy
of the piezo
Potential energy
of tip-surface interaction
dissipation energy from the damping
variables of the system

## Slide 44
V. Dynamic mode
We can rewrite the action as 3 different parts :
where the action is calculated over a period of oscillation
dt
dt
dt
Principle of least action tell us :

## Slide 45
V. Dynamic mode
Equation becomes:
Pure attractive interaction
Pure repulsive interaction
ks is contact spring constant
t time of contact

## Slide 46
V. Dynamic mode
Resonance of the mode
Pure attractive interaction
Pure repulsive interaction

## Slide 47
V. Dynamic mode
Resonance behavior in function of the tip-surface distance
Attractive interaction
Repulsive interaction

## Slide 48
V. Dynamic mode
E) Approach-retract curve
approach
retract
x
Z(t)
Piezo drive
w
Signal generator
piezoZ
Fint
Dz
For a fixed w (slightly off the resonance), analysis of Amplitude A(w) and phase j(w) of the cantilever during the approach (increase of Dz) and retract (decrease of Dz).
Z

## Slide 49
Amplitude of cantilever
Phase of the cantilever
V. Dynamic mode
Attractive interaction
surface

## Slide 50
V. Dynamic mode
Repulsive interaction
surface
slope=1
Amplitude of cantilever
Phase of the cantilever

## Slide 51
Modelling the amplitude and phase for different surface contact
Amplitude
Phase
V. Dynamic mode
k=300
k=0.3
k=0.03
k=0.003

## Slide 52
IV. Dynamic mode
F) Imaging mode
PiezoZ actuator : surface approach
PiezoXY actuators : scan of the surface in XY plane
Surface scanning:
Approach
Scanning X and Y
Top view
Y
X
x
PiezoZ
Dz
PiezoXY
Piezo drive
w
Z(t)
Signal generator

## Slide 53
Constant amplitude imaging (tapping)
V. Dynamic mode
4 quadrants
Laser diode
sample
Setpoint (Aw/A0)
piezoXYZ
PI controller
lockin

## Slide 54
IV. Dynamic mode
Tapping mode operation :
1) Use a cantilever for tapping mode with the first mode frequency around 70kHz-400kHz.
2) Detecting of the resonance of mode. Measurement of the Amplitude in function of the frequency of the piezo driver.
3) Selecting the working frequency (Afree)
4) Selecting the setpoint for imaging (Aw/A0)
5) Approaching the surface until reaching the setpoint (Aw/A0)
6) Scan in X and Y axis keeping the amplitude constant (setpoint) recording the variation of Vz as the topography image

## Slide 55
IV. Dynamic mode
Tapping mode imaging
Afree
PiezoZ
PiezoXY
Piezo drive
w
Aw
PiezoZ
PiezoXY
Piezo drive
w
A
PiezoZ
PiezoXY
Piezo drive
w
Aw
PiezoZ
PiezoXY
Piezo drive
w
Topography image (Vz)

## Slide 56
ADN (100nm x 100 nm)
IV. Dynamic mode
Image AFM en mode dynamique

## Slide 57
VI. Resolution
Tip effect example

## Slide 58
VI. Resolution
A) Vertical sensitivity
Goes to 30pm to 300pm, strongly dependent on the AFM system,
vibration isolation table and environment.
B) Lateral resolution
Strongly dependent on the AFM system and the tip radius:
can go from sub-nanometer up to 10 nm resolution

## Slide 59
OTHER AFM OPERATING MODES

## Slide 60
AFM in fluid
Yeast spore into a membrane (10µm x 10µm)
Force-Volume analysis
Blood cells (tapping)
Contact and tapping mode
Other AFM operating modes

## Slide 61
Other AFM operating modes
Identification and location of a protein on bacterial membrane
@ DOI 10.1002/sca.20177

## Slide 62
Figure 2. Imaging the dynamic clustering of cell-adhesion proteins on the surface of a living cell. (A) Single C. albicans adhesins were localized and stretched on different locations of a single yeast cell using an AFM tip bearing specific antibodies. (B) Molecular recognition map record- ed on a cell that was never probed, thus not subjected to force. Coloured pixels document the detection of single proteins (the blue and red pixels correspond to forces in the 0–150 pN and 150–300 pN range, respectively). Most proteins were isolated and evenly distributed, without any clear evidence for clustering. (C) Subsequent recognition map recorded on a remote area of the same cell following application of force. Proteins in the second map were no longer isolated but clustered into nanodomains referred to as ‘nanoadhesomes’. Reprinted with permission from Alsteens, D. et al. (2010).
Force-volume imaging
Other AFM operating modes
@Current Biology vol21 n6 R212, D.D Müller, Y.F. Dufrêne

## Slide 63
High-resolution force–volume AFM of the extracellular PM surface. a) Topography. Inset of (a) shows three-fold symmetrized cross-correlation average (n = 200) of the bacteriorhodopsin (BR) trimer. Outlined BR shapes (white lines with the seven transmembrane a-helices A–G indicated) were obtained from X-ray and electron crystallographic analyses.[15b] Full color scales correspond to vertical ranges of a) 1.0 nm. Scale bars, 10 nm (raw data) and 2 nm (insets).
Other AFM operating mode
Force-volume imaging
@ DOI: 10.1002/anie.201103991

## Slide 64
Conformational changes of proteins (streptavidin) and DNA observed in buffer solution. (a) The hexagonally packed intermediate (HPI) layer consists of hexameric units having a central core and connecting arms. At the inner surface, the cores exhibit a central pore that occurs in‘open’ (unplugged) and ‘closed’ (plugged) conformations. Repeated imaging of the same surface at 4-min intervals showed that some pores changed their conformation. This conformational change was fully reversible and could be observed over long time periods. The defective pore at the bottom left corner is used as a reference to align the scans. The images were recorded in 100 mM KCl, 20 mM
MgCl2, 10 mM Tris, pH 8.2, using the contact mode. Bar, 15 nm. Shades correpond to a height range of 6 nm. (b) Atomic force microscope images of plasmid DNA. This plasmid contains a 106-bp inverted repeat, which is stabilized in the cruciform conformation by negative DNA supercoiling. Arrows indicate the cruciforms. The frames from the left to right are four repetitive scans over the same area. The hairpin arms of the indicated cruciform were initially quite far away from each other (with an angle of ~1008), but then they moved closer and finally adopted a parallel orientation. The scanning was performed in Tris–EDTA buffer with a NanoScope III MultiMode system (Digital Instruments, Inc., Santa Barbara, CA, USA) operating in the tapping mode. Bar, 40 nm.
Other AFM operating modes
@trends in CELL BIOLOGY (Vol. 9) February 1999
Tapping mode in fluid

## Slide 65
Fluorescence imaging and AFM
Other AFM operating modes

## Slide 66
Nanoindentation (15 nm)
in polymer film
(800nm x 800 nm)
Gold nanoparticles manipulation (8nm)
Nanolithography and nanomanipulation
Other AFM operating modes

## Slide 67
Other AFM operating modes
Nanomechanical imaging
Tapping phase imaging of copolymer
@https://doi.org/10.1021/la0005098

## Slide 68
Other AFM operating modes
PeakForce Tapping
@https://doi.org/10.1021/acs.langmuir.5b02595

## Slide 69
EFM images for 2 inverted polarisation
(15µm x 15 µm)
Transistor default detected by EFM
(80µm x 80 µm)
Applications in nanoelectronics
EFM (Electric Force Microscope)
Other AFM operating modes

## Slide 70
MFM (Magnetic Force Microscope)
Hard disk surface
topography
MFM image
Double pass imaging
Other AFM operating modes

## Slide 71
SThM (Scanning Thermal Microscope)
AFM thermal probe
topography
Thermal conductivity image
Sample of glass fibers into epoxy matrix.
Other AFM operating modes

## Slide 72
Coupling AFM/STM in high vacuum
Other AFM operating modes
@H. Le Sueur. Cryogenic AFM-STM for mesoscopic physics. Condensed Matter. Université Pierre et Marie Curie - Paris VI, 2007.

## Slide 73
AFM-IR : coupling AFM with IR spectroscopy
IR laser
Other AFM operating modes

## Slide 74
Accumulation of biopolymer or lipids
Localisation and quantification
Local composition, TAG, DAG, MAG and FFA differentiation
MICRO-ORGANISMS
Nanoparticules and cell: macrophage
Calcification in human tissues
Extracellular vesicles
Penetration of nanocarriers
Fine structure of the hair…
TISSUE – Human cells
Protein assemblies
Collagen fibrils denaturation
System complex: Collagen-antibiotic
NANOMETRIC SCALE
Bacterial amyloids
Beta structure of amyloids
Prion, lipids bilayer
Field of applications - Biology

## Slide 75
- IR signatures: heterogeneities in ancient tissues or violin sections
Heritage sciences
Polymers sciences
- Polymeric Nps
Nanoparticles
Field of applications
- Investigate parchments degradation
Trace of adjuvant blooming
- Investigation of organic matter in micrometeorites
Astrochemistry
Multilayers:
Structure-cristallinity
Mathurin J., 10.1039/C8AN01239C, Analyst, 2018
A Dazzi, International journal of pharmaceutics Volume 484, Issues 1–2, 2015
J. Mathurin, A&A, 2019
G.Latour, Scientific Report, 2016
A Dazzi, Chem Rev, 2016
Notes: Mettre les références des différentes images
Beside